## Step 1 - Budget
Function: | AC Voltage Measurement
Description: | Fluke Digital Multimeter 189/ SN: 85500034 / Lab ID: E-RMS080
Range: | 6 - 330 V, 60Hz- 20kHz
TEST POINT
Test-Point | 33V @ 60 Hz
Unit | Volts (AC)
Uncertainty Analysis: CMC (Calibration and Measurement Capability)
| Uncertainty Components | Sensitivity Coefficient | Value | Unit | Type | Distribution | Divisor | Std Uncertainty | Degrees of Freedom | Significance Check
|  | (ci) | (xi) |  |  |  |  | u(xi) | υ |  | (u^4)/v | Type | Distribution | Divisor | k-Factor
1 | Repeatability | 1 | 0.001794 | Volts (AC) | A | Normal 1s | 1 | 0.001794 | 58 | 1.534% | 1.7859e-13 | A | Normal 1s | 1 | k=2
2 | Reproducibility | 1 | 0.000471 | Volts (AC) | A | Normal 1s | 1 | 0.000471 | 1 | 0.1057% | 4.9213e-14 | B | Normal 2s | 2 | Student t
3 | Bias | 1 | 0.02 | Volts (AC) | B | Rectangular √3 | 1.7321 | 0.011547 | 1 | 63.54% | 1.7778e-08 |  | Rectangular √3 | 1.7321
4 | Resolution | 1 | 0.01 | Volts (AC) | B | Rectangular √3 | 1.7321 | 0.0057735 | 1e+100 | 15.88%
5 | Drift | 1 | 0.0048961 | Volts (AC) | B | Rectangular √3 | 1.7321 | 0.0028268 | 2 | 3.808% | 3.1925e-11 |  | Rectangular √12 | 3.4641
6 | Reference of Standard Uncertainty | 1 | 0.0086 | Volts (AC) | B | Normal 2s | 2 | 0.0043 | 99 | 8.811% | 3.4533e-12 |  | U-Shaped | 1.4142
7 | Battery Voltage | 1 | 0.001467 | Volts (AC) | A | Normal 1s | 1 | 0.001467 | 58 | 1.026%
8 | Frequency Response | 1 | 0.0057735 | Volts (AC) | B | Rectangular √3 | 1.7321 | 0.0033333 | 1e+100 | 5.295% | 1.2346e-110 |  | Rayleigh | 2.4477
9 |  |  |  | Volts (AC) | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% | 0
10 |  |  |  | Volts (AC) | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% | 0
11 |  |  |  | Volts (AC) | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% | 0
12 |  |  |  | Volts (AC) | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% | 0
13 |  |  |  | Volts (AC) | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% | 0
14 |  |  |  | Volts (AC) | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% | 0
15 |  |  |  | Volts (AC) | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% | 0
16 |  |  |  | Volts (AC) | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% | 0
17 |  |  |  | Volts (AC) | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% | 0
[8 empty] | uc(y) | υeff | Σ
Combined Uncertainty (RSS method) | [7 empty] | 0.014486 | 2 | 100%
Expansion Coefficienct (k) |  | k=2 | [5 empty] | 2
Expanded Uncertainty [ kuc(y) ] | [7 empty] | 0.029 | Volts (AC)
Relative Expanded Uncertainty [ kuc(y) ] | [7 empty] | 0.087879 | %
NOTES
1 | Enter uncertainty component notes here.
2 | Enter uncertainty component notes here.
3 | Enter uncertainty component notes here.
4 | Enter uncertainty component notes here.
5 | Enter uncertainty component notes here.
6 | Enter uncertainty component notes here.
7 | Enter uncertainty component notes here.
8 | Enter uncertainty component notes here.
9 | Enter uncertainty component notes here.
10 | Enter uncertainty component notes here.
11 | Enter uncertainty component notes here.
12 | Enter uncertainty component notes here.
13 | Enter uncertainty component notes here.
14 | Enter uncertainty component notes here.
15 | Enter uncertainty component notes here.
16 | Enter uncertainty component notes here.
17 | Enter uncertainty component notes here.
18 | Enter uncertainty component notes here.
19 | Enter uncertainty component notes here.
20 | Enter uncertainty component notes here.
INSTRUCTIONS
1 | Enter the name of the measurement function or test in B1.*
2 | Enter the description of the measurement equipment or method in B2.*
3 | Enter the measurement range in B3.*
4 | Enter the value of the test point and unit of measurement in B7 and B8.*
5 | Enter the name of the uncertainty components in B14 to B33.
6 | Enter the value of each sensitivity coefficient in C14 to C33. If none used, enter a value of 1 or 0.
7 | Enter the value of each uncertainty component in D14 to D33.
8 | Enter the unit of measure in E14 to E33.*
9 | Enter each uncertainty type (A or B) in F14 to F33.
10 | Select the probability distribution from the dropdown menus in G14 to G33.
11 | Enter the degrees of freedom in J14 to J33.
12 | Select the expansion factor method using the dropdown menus in cells C36.
13 | Enter notes for each uncertainty component in B42 to B61.
*Results may be automatically populated using data from other worksheets. Review before overwriting data

## Frequency Response and Resistan
Data from last 3 Calibration Docs
| 2019 | 2020 | 2021 | Stdev
Frequency Response @ 60VAC 99.95kHz | 99.95 | 99.94 | 99.95 | 0.0057735
Resistance @ 50Mohm | 30.01 | 30 | 30.01 | 0.0057735

## Ref Unc and Ref Stnd
Year | 2022-07-11 | 2021-07-15 | 2020-07-08
| 0.0086 | 0.0086 | 0.0086 | RSU | 0.0086 | RFS | 0
[6 empty] | DOF | 2

## Drift
Year | 2021-07-11 | 2020-07-15 | 2019-07-08 |  | Drift 1 | 0
| 32.98 | 32.98 | 32.97 |  | Drift 2 | 2.681e-05
[5 empty] | ave drift (daily) | 1.3405e-05
[5 empty] | ave drift (annual) | 0.0048961
|  |  |  | 1 step | 0.0048961

## Bias
Year | 2022 | 2021 | 2020 | Bias | -0.02
STD | 33 | 33 | 33
UUT | 32.98 | 32.98 | 32.98

## Freq Respnse and Resistance
*Info was taken from 3 most recent calibration documents
| 2020 | 2021 | 2022 | Stdev
Resistance @ 50Mohm | 30 | 30.01 | 30 | 0.0057735
Frequency Response @ 60VAC 99.95kHz | 99.94 | 99.95 | 99.95 | 0.0057735

## Repeatability & Reproducibility
| Test 1 | Test 2 | Test 1 |  |  | *Regulated power supply was used
1 | 33.17 | 33.17 | Mean | 33.1693 |  | Equipment ID: C-DCP039
2 | 33.17 | 33.17 | StDev | 0.0025371
3 | 33.17 | 33.17 | Dof | 29 |  | *test 1 was performed by me, test 2 was another tech
4 | 33.17 | 33.17 | SS | 0.00018667
5 | 33.17 | 33.17
6 | 33.17 | 33.17 | Test 2
7 | 33.17 | 33.17 | Mean | 33.17
8 | 33.17 | 33.17 | StDev | 1.4454e-14
9 | 33.17 | 33.17 | Dof | 29
10 | 33.17 | 33.17 | SS | 6.0585e-27
11 | 33.16 | 33.17
12 | 33.17 | 33.17 | Total SS | 0.00018667
13 | 33.17 | 33.17 | Total dof | 58
14 | 33.17 | 33.17
15 | 33.17 | 33.17 |  | Unc | DOF
16 | 33.17 | 33.17 | RPT | 0.001794 | 58
17 | 33.17 | 33.17 | RPD | 0.0004714 | 1
18 | 33.16 | 33.17
19 | 33.17 | 33.17
20 | 33.17 | 33.17
21 | 33.17 | 33.17
22 | 33.17 | 33.17
23 | 33.17 | 33.17
24 | 33.17 | 33.17
25 | 33.17 | 33.17
26 | 33.17 | 33.17
27 | 33.17 | 33.17
28 | 33.17 | 33.17
29 | 33.17 | 33.17
30 | 33.17 | 33.17

## Full and Half Battery
| Full Battery | Half Battery | Test 1 |  |  | *Regulated power supply was used
1 | 33.051 | 33.053 | Mean | 33.0518 |  | Equipment ID: C-DCP039
2 | 33.051 | 33.053 | StDev | 0.00046113
3 | 33.051 | 33.053 | Dof | 29 |  | *test 1 was performed with full battery, test 2 was half battery
4 | 33.051 | 33.053 | SS | 6.1667e-06
5 | 33.051 | 33.053
6 | 33.052 | 33.053 | Test 2
7 | 33.052 | 33.053 | Mean | 33.0533
8 | 33.051 | 33.053 | StDev | 0.00046609
9 | 33.052 | 33.053 | Dof | 29
10 | 33.052 | 33.053 | SS | 6.3e-06
11 | 33.052 | 33.053
12 | 33.052 | 33.053 | Total SS | 1.2467e-05
13 | 33.052 | 33.054 | Total dof | 58
14 | 33.052 | 33.053
15 | 33.052 | 33.053 |  | Unc | DOF
16 | 33.052 | 33.053 | RPT | 0.00046362 | 58
17 | 33.052 | 33.053 | RPD | 0.0010371 | 1
18 | 33.052 | 33.053
19 | 33.052 | 33.053
20 | 33.052 | 33.053
21 | 33.052 | 33.053
22 | 33.052 | 33.053
23 | 33.052 | 33.054
24 | 33.052 | 33.054
25 | 33.052 | 33.054
26 | 33.053 | 33.054
27 | 33.052 | 33.054
28 | 33.052 | 33.054
29 | 33.052 | 33.054
30 | 33.052 | 33.054

## Changelog
Date | Updates
2017-10-21 | Version 1.0
* Original Issue
* New: single budget uncertainty calculator: 'simple-uncertainty-calculator.xlsx.'
2018-04-18 | Version 2.0
* Change: Increased number of uncertainty components and notes from 15 to 20 items.
* New: Added error handling functions to formulas for standard uncertainty and significance check cells.
* Change: single budget uncertainty calculator: 'simple-uncertainty-calculator-v2.0.xlsx.'
2019-02-27 | Version 3.0
* Change: Updated standard uncertainty formula.
* Change: Updated probability distribution names.
* New: Added Normal 3s (i.e. 99% C.I.)to probability distributions list
* New: Added error handling functions to formulas for combined uncertainty, coverage factor, expanded uncertainty, effective degrees of freedom, significance check sum,
* Change: single budget uncertainty calculator: 'simple-uncertainty-calculator-v3.0.xlsx.'
2019-03-31 | Version 4.0
* Change: one budget uncertainty calculator: 'simple-uncertainty-calculator-v4-0-one-budget.xlsx.'
* Change: Updated Title, description, range, and test-point section and formatting.
* Change: Updated Notes section and formatting.
* Change: Updated Expanded Uncertainty cell to automatically round result to two significant figures.
* Change: Updated instructions.
* Change: Updated uncertainty component items 9 to 20 to blank cells.
* New: Added summation symbol to sum of significance check values.
* New: Added header with automated file name and revision date.
* New: Added footer section with automated page numbering.
* New: Added dropdown menu to select coverage factor; Student's T Table or k=2.
* New: Added two budget uncertainty calculator with CMC Uncertainty Calculation: 'simple-uncertainty-calculator-v4-0-two-budget.xlsx.'
2019-12-09 | Version 4.1
* Change: Updated Effective Degrees of Freedom function.
* Change: Fixed Uncertainty Type dropdown menu.
* Change: one budget uncertainty calculator title: 'simple-uncertainty-calculator-v4-1-one-budget.xlsx.'
* Change: two budget uncertainty calculator title: 'simple-uncertainty-calculator-v4-1-two-budget.xlsx.'
2020-05-18 | Version 4.2
* Change: one budget uncertainty calculator title: 'simple-uncertainty-calculator-v4-2-one-budget.xlsx.'
* Change: two budget uncertainty calculator title: 'simple-uncertainty-calculator-v4-2-two-budget.xlsx.'
* New: Added Log-Normal distribution to available distributions.
* New: Added automatic page formatting to scale printing to one-page width.
2021-04-18 | Version 5.0
* Change: one budget uncertainty calculator title: 'simple-uncertainty-calculator-v5-0-one-budget.xlsx.'
* Change: two budget uncertainty calculator title: 'simple-uncertainty-calculator-v5-0-two-budget.xlsx.'
* Change: Updated formatting, instructions, fill colors, font colors, and hid unused cells.
* Change: Updated Effective Degrees of Freedom calculation to round result to a whole number.
* Change: Updated TINV function probability from 0.05 to 0.0455.
* Change: Removed grid lines from view on all worksheets.
* Change: Removed Guide worksheet.
* Change: Updated B0 coefficient equation to handle negative results without error.
* New: Added conditional formatting to significance check values. **//Easier to find most significant contributors//**
* New: Added Offset Coefficient (B0) Options to CMC worksheet in two budget template.
* New: Added automated CMC Uncertainty equation generator to CMC worksheet in two budget template.
* New: Validated calculator results by comparison to NIST's Uncertainty Budget Table Template calculator, version 16JAN2013.