## Step 1 - Budget
Function: | Force Gauge
Description: | Load Cells Tension and Compression - Morehouse Ultra Precision w/PMC Calibration Stand ID # 02-2027
Range: | Up to 2000 lbf
LOW POINT
Test-Point | 200
Unit | lbf
Uncertainty Analysis: CMC (Calibration and Measurement Capability)
| Uncertainty Components | Sensitivity Coefficient | Value | Unit | Type | Distribution | Divisor | Std Uncertainty | Degrees of Freedom | Significance Check |  |  | Type | Distribution | Divisor | k-Factor
|  | (ci) | (xi) |  |  |  |  | u(xi) | υ |  |  | (u^4)/v | A | Normal 1s | 1 | k=2
1 | Repeatability | 1 | 0.00058695 | lbf | A | Normal 1s | 1 | 0.00058695 | 19 | 2.968% |  | 6.2469e-15 | B | Normal 2s | 2 | Student t
2 | Reproducibility | 1 | 0.0016126 | lbf | A | Normal 1s | 1 | 0.0016126 | 1 | 22.4% |  | 6.7618e-12 |  | Normal 3s | 2.5758
3 | Stability | 1 | 0.0014832 | lbf | B | Normal 1s | 1 | 0.0014832 | 99 | 18.95% |  | 4.8889e-14 |  | Rectangular √3 | 1.7321
4 | Bias | 1 | 0.001 | lbf | B | Normal 1s | 1 | 0.001 | 1 | 8.615% |  | 1e-12 |  | Rectangular √12 | 3.4641
5 | Drift | 1 | 0.0020069 | lbf | B | Normal 1s | 1 | 0.0020069 | 2 | 34.7% |  | 8.1103e-12 |  | Triangular | 2.4495
6 | Resolution | 1 | 0.001 | lbf | B | Rectangular √3 | 1.7321 | 0.00057735 | 1e+100 | 2.872% |  | 1.1111e-113 |  | U-Shaped | 1.4142
7 | Reference Standard Uncertainty | 1 | 0.0021 | lbf | B | Normal 2s | 2 | 0.00105 | 99 | 9.498% |  | 1.2278e-14 |  | Rayleigh | 2.4477
8 | Reference Standard Stability | 1 | 0 | lbf | B | Normal 1s | 1 | 0 | 2 | 0% |  | 0
9 | Reference Standard Drift | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
10 | Reference Standard Resolution | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
11 | Machine Control | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
12 | Zero Stability | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
13 | Creep | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
14 | Axial Alignment | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
15 | Fixture Repeatability | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
16 | Temperature | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
17 | Hysteresis | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
18 | Settling Time / Loading Rate / Dwell | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
19 | Applied Force Setting Error | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
20 | Adapter Geometry | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
[8 empty] | uc(y) | υeff | Σ
Combined Uncertainty (RSS method) | [7 empty] | 0.0034071 | 8.4538 | 100%
Expansion Coefficienct (k) |  | k=2 | [5 empty] | 2
Expanded Uncertainty [ kuc(y) ] | [7 empty] | 0.0068 | lbf
HIGH POINT
Test-Point | 2000
Unit | lbf
Uncertainty Analysis: CMC (Calibration and Measurement Capability)
| Uncertainty Components | Sensitivity Coefficient | Value | Unit | Type | Distribution | Divisor | Std Uncertainty | Degrees of Freedom | Significance Check
|  | (ci) | (xi) |  |  |  |  | u(xi) | υ |  |  | (u^4)/v
1 | Repeatability | 1 | 0.0051474 | lbf | A | Normal 1s | 1 | 0.0051474 | 19 | 9.274% |  | 3.6948e-11
2 | Reproducibility | 1 | 0.015442 | lbf | A | Normal 1s | 1 | 0.015442 | 1 | 83.46% |  | 5.686e-08
3 | Stability | 1 | 0.0026077 | lbf | B | Normal 1s | 1 | 0.0026077 | 99 | 2.38% |  | 4.6707e-13
4 | Bias | 1 | -0.001 | lbf | B | Normal 1s | 1 | -0.001 | 1 | 0.35% |  | 1e-12
5 | Drift | 1 | 0.0030328 | lbf | B | Normal 1s | 1 | 0.0030328 | 2 | 3.219% |  | 4.2303e-11
6 | Resolution | 1 | 0.001 | lbf | B | Rectangular √3 | 1.7321 | 0.00057735 | 1e+100 | 0.1167% |  | 1.1111e-113
7 | Reference Standard Uncertainty | 1 | 0.0037 | lbf | B | Normal 2s | 2 | 0.00185 | 99 | 1.198% |  | 1.1832e-13
8 | Reference Standard Stability | 1 | 4.8487e-19 | lbf | B | Normal 1s | 1 | 4.8487e-19 | 2 | 8.229e-32% |  | 2.7636e-74
9 | Reference Standard Drift | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
10 | Reference Standard Resolution | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
11 | Machine Control | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
12 | Zero Stability | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
13 | Creep | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
14 | Axial Alignment | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
15 | Fixture Repeatability | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
16 | Temperature | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
17 | Hysteresis | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
18 | Settling Time / Loading Rate / Dwell | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
19 | Applied Force Setting Error | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
20 | Adapter Geometry | 1 | 0 | lbf | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
[8 empty] | uc(y) | υeff | Σ
Combined Uncertainty (RSS method) | [7 empty] | 0.016903 | 1.4335 | 100%
Expansion Coefficienct (k) |  | k=2 | [5 empty] | 2
Expanded Uncertainty [ kuc(y) ] | [7 empty] | 0.034 | lbf
NOTES | [11 empty] | INSTRUCTIONS
1 | Enter uncertainty component notes here. | [9 empty] | 1 | Enter the name of the measurement function or test in B1.*
2 | Enter uncertainty component notes here. | [9 empty] | 2 | Enter the description of the measurement equipment or method in B2.*
3 | Enter uncertainty component notes here. | [9 empty] | 3 | Enter the value measurement range in B3.*
4 | Enter uncertainty component notes here. | [9 empty] | 4 | Enter the value of the test point and unit of measurement in B7 and B8, and B41 and B42.*
5 | Enter uncertainty component notes here. | [9 empty] | 5 | Enter the name of the uncertainty components in B14 to B33 and B48 to B67.
6 | Enter uncertainty component notes here. | [9 empty] | 6 | Enter the value of each sensitivity coefficient in C14 to C33 and C48 to C67. If none used, enter a value of 1.
7 | Enter uncertainty component notes here. | [9 empty] | 7 | Enter the value of each uncertainty component in D14 to D33 and D48 to D67.
8 | Enter uncertainty component notes here. | [9 empty] | 8 | Enter the unit of measure in E14 to E33 and E48 to E67.*
9 | Enter uncertainty component notes here. | [9 empty] | 9 | Enter each uncertainty type (A or B) in F14 to F33 and F48 to F67.
10 | Enter uncertainty component notes here. | [9 empty] | 10 | Select the probability distribution from the dropdown menus in G14 to G33 and G48 to G67.
11 | Enter uncertainty component notes here. | [9 empty] | 11 | Enter the degrees of freedom in J14 to J33 and J48 to J67.
12 | Enter uncertainty component notes here. | [9 empty] | 12 | Select the expansion factor method using the dropdown menus in cells B36 and B70.
13 | Enter uncertainty component notes here. | [9 empty] | 13 | Enter notes for each uncertainty component in B76 to B95.
14 | Enter uncertainty component notes here. | [10 empty] | *Results may be automatically populated using data from other worksheets. Review before overwriting data
15 | Enter uncertainty component notes here.
16 | Enter uncertainty component notes here.
17 | Enter uncertainty component notes here.
18 | Enter uncertainty component notes here.
19 | Enter uncertainty component notes here.
20 | Enter uncertainty component notes here.

## Step 2 - CMC
CMC UNCERTAINTY
INTERPOLATION |  |  |  |  | COEFFICIENTS |  |  |  | GAIN COEFFICIENT (B1) OPTIONS
| X | Unit | Y | Unit | Gain Coefficient | B1 | 1.5e-05 |  | Percent | 0.0015 | %
Low | 200 | lbf | 0.0068 | lbf | Offset Coefficient | B0 | 0.0038 | lbf | PPM | 15 | ppm
High | 2000 | lbf | 0.034 | lbf
δ | 1800 | lbf | 0.0272 | lbf
VERIFICATION | [6 empty] | INSTRUCTIONS
# | X | Unit | Y | Unit |  | 1 | Enter the Low test-point value in cell B5.
1 | 1 | lbf | 0.0038 | lbf |  | 2 | Enter the Low test-point unit of measurement in cell C5.
2 | 2 | lbf | 0.0038 | lbf |  | 3 | Enter the High test-point value in cell B6.
3 | 3 | lbf | 0.0038 | lbf |  | 4 | Determine your CMC Uncertainty equation:
4 | 4 | lbf | 0.0039 | lbf |  |  | a | If B1 is zero and B0 is not, your CMC Uncertainty is B0
5 | 5 | lbf | 0.0039 | lbf |  |  | b | If B1 is not zero and B0 is zero, you CMC Uncertainty is B1
6 | 6 | lbf | 0.0039 | lbf |  |  | c | If B1 is not zero and B0 is not zero, your CMC Uncertainty is B1 + B0
7 | 7 | lbf | 0.0039 | lbf |  |  | d | If you do not want to use a CMC Uncertainty equation, select the largest uncertainty value for the measurement range.
8 | 8 | lbf | 0.0039 | lbf |  | 5 | For verification, enter your desired test-points (including the test-points from your analysis) in cells B13 through B22.
9 | 9 | lbf | 0.0039 | lbf
10 | 10 | lbf | 0.004 | lbf
𝑈_𝐶𝑀𝐶=𝛽_1 𝑥_𝑖+𝛽_0

## Guide
Name of Measurement Function or Test
Description of Measurement Equipment or Method
Range: [Value & Unit] to [Value & Unit]
Test Point: [Value & Unit]
Uncertainty Analysis: CMC (Calibration and Measurement Capability)
| Uncertainty Components | Sensitivity Coefficient | Value | Unit | Type | Distribution | Divisor | Std Uncertainty | Degrees of Freedom | Significance Check |  |  | Type | Distribution | Divisor
|  | (ci) | (xi) |  |  |  |  | u(xi) | υ |  |  | (c^4*u^4)/v | A | Normal 1 | 1
1 | Repeatability | 1 | 0.000102 | kPa | A | Normal 1 | 1 | 0.000102 | 19 | 0.04114% |  | 5.697e-18 | B | Normal 2 | 2
2 | Reproducibility | 1 | 0.000174 | kPa | A | Normal 1 | 1 | 0.000174 | 3 | 0.1197% |  | 3.0555e-16 |  | Rectangular √3 | 1.7321
3 | Stability | 1 | 0.005 | kPa | B | Normal 1 | 1 | 0.005 | 99 | 98.85% |  | 6.3131e-12 |  | Rectangular √12 | 3.4641
4 | Bias | 1 | -0.0005 | kPa | B | Normal 1 | 1 | -0.0005 | 1 | 0.9885% |  | 6.25e-14 |  | Triangular | 2.4495
5 | Drift | 1 | 0 | kPa | B | Normal 1 | 1 | 0 | 2 | 0% |  | 0 |  | U-Shaped | 1.4142
6 | Resolution | 1 | 1e-05 | kPa | B | Rectangular √3 | 1.7321 | 5.7735e-06 | 1e+100 | 0.0001318% |  | 1e-120 |  | Rayleigh | 2.4477
7 | Reference Standard Uncertainty | 1 | 6.1e-05 | kPa | B | Normal 2 | 2 | 3.05e-05 | 99 | 0.003678% |  | 1.3986e-19
8 | Reference Standard Stability | 1 | 0 | kPa | B | Normal 1 | 1 | 0 | 2 | 0% |  | 0
9 | Other Factor 1 | 1 | 0 | kPa | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
10 | Other Factor 2 | 1 | 0 | kPa | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
11 | Other Factor 3 | 1 | 0 | kPa | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
12 | Other Factor 4 | 1 | 0 | kPa | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
13 | Other Factor 5 | 1 | 0 | kPa | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
14 | Other Factor 6 | 1 | 0 | kPa | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
15 | Other Factor 7 | 1 | 0 | kPa | B | Rectangular √3 | 1.7321 | 0 | 1e+100 | 0% |  | 0
[8 empty] | uc(y) | υeff
Combined Uncertainty (RSS method) | [7 empty] | 0.0050291 | 100.3251 | 100%
Expansion Coefficienct (k) | [7 empty] | 1.984
Expanded Uncertainty [ kuc(y) ] | [7 empty] | 0.0099776 | kPa
NOTES | [11 empty] | INSTRUCTIONS
1 | Uncertainty component notes | [9 empty] | 1 | Enter the name of the measurement function or test in A1.
2 | Uncertainty component notes | [9 empty] | 2 | Enter the description of the measurement equipment or method in A2.
3 | Uncertainty component notes | [9 empty] | 3 | Enter the value measurement range in A3.
4 | Uncertainty component notes | [9 empty] | 4 | Enter the value of the test point in A4
5 | Uncertainty component notes | [9 empty] | 5 | Enter the name of the uncertainty components in B10 to B24.
6 | Uncertainty component notes | [9 empty] | 6 | Enter the value of each sensitivity coefficient in C10 to C24. If none used, enter a value of 1.
7 | Uncertainty component notes | [9 empty] | 7 | Enter the value of each uncertainty component in D10 to D24.
8 | Uncertainty component notes | [9 empty] | 8 | Enter the unit of measure in E10 to E24.
9 | Uncertainty component notes | [9 empty] | 9 | Enter each uncertainty type (A or B) in F10 to F24.
10 | Uncertainty component notes | [9 empty] | 10 | Select the probability distribution from the dropdown menu in G10 to G24.
11 | Uncertainty component notes | [9 empty] | 11 | Enter the degrees of freedom in J10 to J24.
12 | Uncertainty component notes | [9 empty] | 12 | Enter notes for each uncertainty component in B33 to B47.
13 | Uncertainty component notes
14 | Uncertainty component notes
15 | Uncertainty component notes
LEGEND
Type to enter description of the measurement function, measurement method or equipment, measurement range, and test point
Type to enter uncertainty component name
Type to enter value of sensitivity coefficient
Type to enter value of uncertainty component
Type to enter unit of measure
Type to enter uncertainty type: A or B
Use dropdown menu to select probability distribution
Type to enter value of degrees of freedom
Type to enter notes about uncertainty component

## Changelog
Date | Updates
2018-02-28 | Added error handling to: standard uncertainty, significance check, and weighted variance.
Added dropdown lists for Uncertainty type and probability distribution.
2019-02-27 | Increased the calculator to handle 20 sources of uncertainty
Added error handling to: combined uncertainty, expansion factor, expanded uncertainty, and effective degrees of freedom cells.
Updated Normal distributions to include the letter "s" for sigma.
Added Normal 3s distribution for 3-sigma or 99% confidence intervals.
Updated Notes text to: "Enter uncertainty component notes here."
2019-03-22 | Updated Measurement function, desctiption, range, and test-points cells for easier user entry.
Added a two budget option with CMC Uncertainty calculator.
Updated expanded uncertainty formula to automatically round to two significant figures.
Added dropdown list to allow users to switch between k=2 and student's t tables.
Added automatic file name printed in the header.
Added automatic date and page numbering to the footer.
2019-03-27 | Added Survey worksheet.
Added Cal Data worksheet.
Added R&R worksheet.
Added Linearity worksheet.
Added automatic calculation and linking of results to uncertainty in Budget worksheet for factors 1 through 8.
2019-12-09 | Updated Effective Degrees of Freedom calculation in 'Step 1 - Budget.'
Fixed Uncertainty Type dropdown menu in 'Step 1 - Budget.'